★ EXAMPLE – FOR REFERENCE ONLY ★

CORRECTIVE AND PREVENTIVE ACTION (CAPA)

Form

| CAPA Number: CAPA-2025-017 | Date Initiated: 2025-01-08 |
| --- | --- |
| CAPA Owner: Dr. Maria Schmidt, Quality Manager | Target Completion Date: 2025-03-15 |

| Source: Internal Audit External Audit Customer Complaint Nonconformance PMS/Vigilance Management Review Supplier Issue Process Deviation Other: _______________ | Priority / Risk Level: Critical (Patient Safety Impact) High (Regulatory/Major Quality) Medium (Moderate Impact) Low (Minor / Improvement) |
| --- | --- |

1. PROBLEM DESCRIPTION [ISO 13485 §8.5.2a / FDA 820.100(a)(1)]

| Problem Statement | On 2025-01-06, customer complaint CC-2025-0042 was received from St. Anna Hospital. The complaint reported that 47 units of sterile urinary catheters (REF UC-200-16FR) from lot #L2024-1187 displayed an incorrect expiration date on the primary packaging. Label showed: EXP 2025-06 (incorrect) Correct date should be: EXP 2027-06 The shortened expiration date would cause premature disposal of compliant product, resulting in unnecessary waste and potential supply issues for the customer. |
| --- | --- |
| Affected Product(s) / Process(es) | Sterile Urinary Catheter, 16FR (REF UC-200-16FR), Lot #L2024-1187, 2,400 units produced |
| Reference Documents | CC-2025-0042, DHR-L2024-1187, NC-2025-0023 |
| Impact Assessment | Patient Safety: LOW – Product sterility and function not affected, only label error Regulatory: HIGH – Mislabeling violates 21 CFR 820.120 and EU MDR Annex I §23 Business: MEDIUM – Customer trust affected, potential recall costs Scope: 2,400 units distributed to 12 customers across EU and US markets |

2. IMMEDIATE CONTAINMENT ACTIONS [FDA 820.100(a)(3)]

| Containment Actions Taken | 1. Production line for UC-200 series stopped immediately (2025-01-06, 14:30) 2. Remaining inventory of lot L2024-1187 quarantined (312 units in warehouse) 3. Customer notification sent to all 12 affected customers within 24 hours 4. Advisory issued: Product safe to use, apply corrective label over expiry date 5. Corrective labels printed and shipped to customers (2025-01-08) 6. No field safety corrective action (recall) required per risk assessment RA-2025-008 |
| --- | --- |
| Containment Verified By | Thomas Weber, Production Manager – 2025-01-08 |

3. ROOT CAUSE INVESTIGATION [ISO 13485 §8.5.2b / FDA 820.100(a)(2)]

| RCA Method Used | 5-Why Analysis Fishbone / Ishikawa Diagram Fault Tree Analysis Pareto Analysis Other: _______________ |
| --- | --- |
| Investigation Team | Dr. M. Schmidt (QM), T. Weber (Prod), K. Hoffmann (Labeling), S. Braun (QC) |

5-Why Analysis (Fishbone diagram attached as Appendix A)

Why 1: Why was the wrong expiration date printed? → The label template contained an incorrect shelf life parameter (24 months instead of 48 months)

Why 2: Why did the template have the wrong shelf life? → Template was copied from UC-100 series (24-month shelf life) when creating UC-200 series

Why 3: Why wasn't the shelf life parameter updated after copying? → No checklist exists for verifying all parameters when creating new label templates

Why 4: Why wasn't the error caught during label approval? → Label approval SOP focuses on text/graphics review, no explicit verification against DHF/DMR data

Why 5: Why doesn't the SOP require DHF/DMR verification? → SOP-LAB-002 was written in 2018 and has not been updated to include this critical check (ROOT CAUSE)

| Root Cause Statement | ROOT CAUSE: SOP-LAB-002 "Label Design and Approval" lacks a mandatory verification step requiring cross-check of critical label parameters (expiration date, shelf life, storage conditions) against the Device Master Record (DMR) and Design History File (DHF). CONTRIBUTING FACTOR: No standardized template creation checklist exists, allowing copy-paste errors to propagate undetected. |
| --- | --- |
| Evidence / Data Supporting Root Cause | • Review of SOP-LAB-002 Rev 3 confirms no DHF/DMR verification requirement • Interview with K. Hoffmann confirmed template was copied from UC-100 • Label approval record LAB-2024-0892 shows no shelf life verification performed • DMR-UC-200 clearly states 48-month shelf life based on stability study STB-2022-014 • Similar near-miss found in UC-150 template (caught before production) |

4. CORRECTIVE ACTION PLAN [ISO 13485 §8.5.2c-d / FDA 820.100(a)(3-4)]

| # | Corrective Action | Responsible | Due Date | Status |
| --- | --- | --- | --- | --- |
| 1 | Correct UC-200 label template, update shelf life to 48 months | K. Hoffmann | 2025-01-15 | COMPLETE |
| 2 | Revise SOP-LAB-002 to include mandatory DHF/DMR verification step | Dr. M. Schmidt | 2025-02-01 | COMPLETE |
| 3 | Create Label Template Creation Checklist (FRM-LAB-015) | K. Hoffmann | 2025-02-15 | COMPLETE |
| 4 | Train all labeling personnel on revised SOP and checklist | S. Braun | 2025-02-28 | COMPLETE |
| 5 | Audit all active label templates against DMR (32 SKUs) | QC Team | 2025-03-15 | COMPLETE |

5. PREVENTIVE ACTION [ISO 13485 §8.5.3 / FDA 820.100(a)(3)]

| Preventive Actions | 1. Implement automated label parameter validation in ERP system (Project IT-2025-04) - System will flag mismatches between label template and DMR before print approval - Target implementation: Q3 2025 2. Add annual label audit to internal audit schedule (AUD-PLAN-2025 updated) 3. Include labeling process in next Management Review (scheduled 2025-04-15) |
| --- | --- |
| Horizontal Deployment | Other products reviewed for similar risk Other processes reviewed for similar risk Supplier/vendor notification required Not applicable (justify below) |
| Justification / Notes | All 32 active product label templates audited. One additional discrepancy found in UC-150 template (storage temperature) - corrected under NC-2025-0031. |

6. VERIFICATION OF IMPLEMENTATION [ISO 13485 §8.5.2e / FDA 820.100(a)(4-5)]

| Verification Activities | Actions completed as planned Documentation updated (SOPs, WIs, Forms) Training completed and documented Process/Design changes implemented Actions do not adversely affect product safety/performance |
| --- | --- |
| Evidence of Implementation | • SOP-LAB-002 Rev 4 effective 2025-02-01 (ECO-2025-0012) • FRM-LAB-015 "Label Template Creation Checklist" approved 2025-02-15 • Training records TR-2025-0089 through TR-2025-0094 (6 employees) • Label audit report QA-RPT-2025-008 completed 2025-03-15 |
| Verified By / Date | S. Braun, QC Supervisor – 2025-03-15 |

7. EFFECTIVENESS VERIFICATION [ISO 13485 §8.5.2f / FDA 820.100(a)(4)]

| Effectiveness Check Date | 2025-06-15 (90 days after last action completed) |
| --- | --- |
| Effectiveness Criteria | 1. Zero label-related nonconformances in 90-day monitoring period 2. 100% compliance with new checklist (verified by audit of 10 random label approvals) 3. Zero customer complaints related to labeling errors |
| Effectiveness Results | Monitoring Period: 2025-03-16 to 2025-06-15 Results: • Label NCs in period: 0 (vs. 3 in prior 90-day period) ✓ • Checklist compliance audit: 10/10 approvals compliant (100%) ✓ • Customer complaints (labeling): 0 ✓ New label templates created in period: 4 (all verified against DMR per new SOP) |
| Effectiveness Determination | EFFECTIVE - Root cause eliminated, no recurrence PARTIALLY EFFECTIVE - Improvement observed, monitoring continues NOT EFFECTIVE - New CAPA required (Reference: CAPA-____-___) |

8. CAPA CLOSURE [ISO 13485 §8.5.2g / FDA 820.100(a)(6-7)]

| Management Review Reference | MR-2025-Q2, Agenda Item 5.3, Meeting Date: 2025-07-10 |
| --- | --- |
| Lessons Learned | 1. Legacy SOPs require periodic review to ensure they address current risks 2. Copy-paste workflow for templates needs systematic verification controls 3. Automated validation (ERP integration) should be prioritized for critical parameters 4. Cross-functional investigation teams improve root cause identification |

| APPROVALS |  |  |
| --- | --- | --- |
| CAPA Owner | Dr. Maria Schmidt | 2025-06-20 |
| Quality Manager | Dr. Maria Schmidt | 2025-06-20 |
| Management Representative | Hans Müller, CEO | 2025-07-10 |

This form complies with ISO 13485:2016 §8.5.2-8.5.3, FDA 21 CFR 820.100, and EU MDR 2017/745 Annex IX requirements.

Example template provided by QCore Consulting | qcore-consulting.de